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DESCRIPTION:
Plate reader spectrophotometer that has 340–850 nm wavelength range, enabling a wide variety of applications from enzyme kinetic studies to Lowry assays. The instrument provides fast and accurate measurements enabling complete 96-well plate reading in less than seven seconds. The instrument is also equipped with linear shaking as standard. A model with an incubator enables incubation temperatures up to 50°C and the capacity to read 384-well plates. software supports optimal use of the instrument with a visual workflow, effortless data analysis and exporting capabilities.

TECHNICAL SPECIFICATIONS:
Description			Plate reader
Warranty			Minimum 1 Year
Detector Type			Photodetector
Light Source			Quartz-halogen lamp
Reading Options		96 Well Plate; 384 Well Plate
Read-out Range		0 to 6 Abs
Software Type			Software with option to upgrade
Plate Type			96-well plates
Wavelength Selection		Filters
Wavelength Range	340 to 850nm (Flexible to upgrade), able to Scan and determine lambda max.
Filters	8-position filter wheel; filters installed: 405nm, 450nm, 540 and   620nm supply with two more filters, 690nm and 790nm (Flexible to installed new filters)
Linearity			0–3 Abs, ± 2% at 405nm, 96-well plate
				384-well plate: 0 to 2.5 Abs, ±2% at 405 nm
Compatibility			Robot
Resolution			0.001 Abs
Accuracy			± 1% (0–3 Abs) or ± 0.003 Abs
Precision			CV ≤ 0.2% (0.3–3 Abs) at 405nm
Incubator Temperature	Up to 50°C
Shaking			Linear (able to shake plate during processing)
Measurement Speed	12 seconds, 96-well plate, fast mode; 13 seconds, 96-well plate, normal mode
Interface			User: On-board or PC control
PC should be provided along with the instrument for software installation.


Note: The instrument should be of country of origin with minimum of 1 year warranty. Only authorized distributor of the particular instrument will be entertained 

